## Slide 2: The Wheatstone Bridge
Figure 4.10
Use Kirchhoff’s Voltage Law:
Ei = I1R1 + I2R2
Ei = I3R3 + I4R4
Eo = I4R4 - I2R2
Eo = -I3R3 + I1R1
Further, I1 = I2 and I3 = I4.
Combining these equations gives

## Slide 3
What is the Wheatstone
bridge equation if another
resistor, Rx, is added in
parallel with R1 ?
Rx
Now 1/Rnew = 1/Rx+1/R1
To solve, simply substitute
Rnew = R1Rx / (R1+Rx)
into the WB equation for R1.

## Slide 4
One advantage of using two resistors in parallel in
one leg of the WB is that the added resistor can be
located remotely from the actual WB, such as in a flow.
Here, that resistor can serve as a sensor.
Rx
FLOW

## Slide 5
When Eo = 0, the WB is ‘balanced’ →
→
When the WB is balanced and 3 of the 4 resistances
are known, the 4th (unknown) resistance can be found
using the balanced WB equation. The is called the null
method.

## Slide 6
Now consider the case when all 4 resistors are the
same initially and, then, one resistance, say R1 is changed
by an amount R. This is called the deflection method.
Often, R is associated with a change in a physical variable.

## Slide 7: Cantilever Beam with Four Strain Gages
When F is applied as shown (downward), R1 and R4
increase by R (due to extension), and R2 and R3
decrease by R (due to compression).
Figure 4.11
From solid mechanics, for a cantilever beam, L or C ~ F
Figure 6.2

## Slide 8: The Cantilever Beam with Four Gages
The bridge equation
This is the voltage/force relation for a cantilever beam.
Because R ~ L or C ~ F, Eo = constant x F
becomes
which reduces to Eo = Ei (R/R).

## Slide 9: In-Class Example
[c] G = amp.output / amp. input
= Em/Eo = 10/0.0357 = 210.11 >> set at 200.
[d] R60 ºC = R20 ºC[1+0.0005(60-20)] = (25)(1.02) = 25.5 
R =R60 ºC – R20 ºC = 25.5 – 25 = 0.5 
Eo = Ei[(R/R)/(4+2R/R)] = 0.1/4.04 = 0.0248 V
Em = G Eo = (200)(0.0248) = 4.95 V

## Slide 10: Loading Error
When measuring a current, the input impedance of the
ammeter must be much less than the equivalent circuit’s
output impedance.
When measuring a voltage, the input impedance of the
voltmeter must be much greater than the equivalent
circuit’s output impedance.